Ventilator Waveform Analysis

A Comprehensive Guide for Advanced Practice Providers

Standalone Reference

UTHealth Houston HVI

# Section 1: Fundamentals

## 1.1 The Three Scalars and Two Loops

The ventilator displays two categories of graphics:

Scalar waveforms plot a variable against time. There are three:

- Pressure-time the workhorse; shows airway pressure throughout the breath cycle

- Flow-time reveals the rate and pattern of gas delivery and exhalation

- Volume-time tracks the cumulative volume delivered and exhaled

Loops plot two variables against each other during a single breath:

- Pressure-volume (PV) loop compliance, overdistension, work of breathing

- Flow-volume (FV) loop airway resistance, air trapping, bronchodilator response

💡 Clinical Pearl: Scalars tell you WHEN something happens in the breath cycle. Loops tell you HOW the lung is behaving mechanically. You need both.

## 1.2 The Dependent vs. Independent Variable

The single most important concept in waveform interpretation is understanding which variables the ventilator controls versus which are dependent on patient mechanics.

| Mode | Ventilator Controls | Patient Mechanics Determine |
| --- | --- | --- |
| ACVC | Tidal volume, flow rate, flow pattern | Peak pressure, plateau pressure |
| ACPC | Inspiratory pressure, inspiratory time | Tidal volume, flow pattern |
| PSV | Pressure support level | Tidal volume, flow, inspiratory time |
| PRVC | Target Vt (via auto-adjusting pressure) | Pressure (breath-to-breath), flow |

The dependent variable is what you monitor — it is the waveform that reveals changes in lung mechanics.

[IMAGE: Side-by-side comparison showing ACVC vs ACPC scalar waveforms (pressure, flow, volume). Label which waveforms are set vs variable in each mode.]

# Section 2: Scalar Waveforms

## 2.1 Pressure-Time Scalar

The pressure-time waveform shows airway pressure (y-axis, in cmH₂O) over time (x-axis, in seconds).

### Normal Appearance — ACVC (Constant Flow)

With constant (square) flow delivery, the pressure waveform rises in a linear ramp during inspiration, reaching peak inspiratory pressure (PIP), then drops during expiration back to PEEP.

Key landmarks: Baseline (PEEP), Rise phase, Peak Inspiratory Pressure (PIP), Plateau Pressure (Pplat), and the PIP–Pplat gradient (reflecting airway resistance).

[IMAGE: Labeled pressure-time scalar in ACVC showing: baseline PEEP, linear rise, PIP, inspiratory hold/plateau, PIP-Pplat gap as airway resistance, Pplat as alveolar pressure, and return to PEEP.]

📋 Key Equations: PIP = Resistive Pressure + Elastic Pressure + PEEP. Resistive Pressure = PIP − Pplat (airway resistance). Elastic Pressure = Pplat − PEEP (driving pressure / compliance).

### What the Pressure Waveform Tells You

| Finding | What It Means | Action |
| --- | --- | --- |
| Rising PIP with stable Pplat | Increased airway resistance (secretions, bronchospasm, kinked ETT) | Suction, bronchodilators, check ETT |
| Rising PIP with rising Pplat | Decreased lung compliance (ARDS, PTX, edema, abdominal distension) | Chest imaging, reduce Vt, increase PEEP |
| Concave scooped pressure (ACVC) | Flow starvation — patient demand exceeds set flow | Increase flow rate or switch to ACPC |
| Pressure spike at end of inspiration | Delayed cycling — vent still inflating after patient wants to exhale | Shorten Ti, increase cycle-off threshold |
| Pressure not returning to set PEEP | Auto-PEEP / air trapping | Extend expiratory time, reduce RR |
| Negative deflection before breath triggers | Patient effort (trigger work) | Adjust trigger sensitivity |

## 2.2 Flow-Time Scalar

The flow-time waveform shows gas flow (L/min) over time. Inspiration is above the baseline; expiration is below.

[IMAGE: Three flow-time waveforms side by side: (1) ACVC square flow, (2) ACVC decelerating flow, (3) ACPC decelerating flow. Label inspiration above baseline, expiration below.]

### What the Flow Waveform Tells You

| Finding | What It Means | Action |
| --- | --- | --- |
| Expiratory flow does not return to zero | Auto-PEEP — air trapping | Extend expiratory time, reduce RR, bronchodilators |
| Abrupt expiratory flow cutoff | Severe air trapping with premature breath delivery | Urgently extend I:E ratio |
| Saw-tooth pattern on expiratory limb | Secretions in the airway | Suction the patient |
| Variable inspiratory flow (ACPC) | Changing compliance or patient effort | Assess for clinical change, monitor Vt |
| Small flow deflections that don’t trigger | Ineffective efforts | Reduce trigger threshold, address auto-PEEP |

[IMAGE: Flow-time scalar showing saw-tooth expiratory pattern from secretions. Label the oscillations on the expiratory limb.]

[IMAGE: Flow-time scalar showing expiratory flow not returning to zero (auto-PEEP). Arrow where next breath starts before flow reaches zero.]

## 2.3 Volume-Time Scalar

The volume-time waveform tracks cumulative volume (mL) over time. It rises during inspiration, then falls during expiration back toward baseline.

| Finding | What It Means | Action |
| --- | --- | --- |
| Exhaled volume doesn’t return to baseline | Circuit leak — cuff leak, chest tube air leak | Check cuff pressure, circuit connections |
| Gradual upward drift of baseline | Progressive air trapping / auto-PEEP | Address auto-PEEP |
| Variable tidal volumes (ACPC) | Changing compliance or variable patient effort | Assess patient; set appropriate Vt alarms |

[IMAGE: Volume-time scalar showing a circuit leak — inspired volume rises to set Vt but exhaled volume doesn’t return to baseline. Label the leak volume gap.]

# Section 3: Waveform Patterns by Mode

## 3.1 ACVC — Assist Control Volume Control

Pressure: Linear ramp up to PIP → plateau if inspiratory hold applied → drop to PEEP

Flow: Square (constant) during inspiration → exponential decay during expiration

Volume: Linear rise → exponential fall

[IMAGE: Complete ACVC scalar set — all three waveforms stacked vertically and time-aligned. Label PIP, Pplat, set flow rate, Vt, PEEP, inspiration, expiration.]

## 3.2 ACPC — Assist Control Pressure Control

Pressure: Rapid rise to set pressure → flat plateau for duration of Ti → drop to PEEP

Flow: Decelerating from peak → reaches zero at end of Ti → passive expiration

Volume: Curved rise (concave down) → exponential fall

[IMAGE: Complete ACPC scalar set — all three waveforms stacked vertically and time-aligned. Label set pressure, Ti, decelerating flow, Vt (variable), PEEP.]

## 3.3 PSV — Pressure Support Ventilation

Pressure: Rapid rise to set PS level above PEEP → maintained until flow drops to cycle-off threshold → drop to PEEP

Flow: Decelerating from peak → terminates when flow reaches cycle-off percentage (typically 25% of peak)

Volume: Curved rise → fall (variable Vt depending on effort and mechanics)

[IMAGE: Complete PSV scalar set. Label PS level, PEEP, cycle-off point on flow waveform, variable Vt.]

## 3.4 PRVC — Pressure Regulated Volume Control

Looks like ACPC breath-to-breath, but the pressure level auto-adjusts to achieve a target Vt. Watch the pressure trend over multiple breaths — if the ventilator is progressively increasing pressure, compliance is worsening.

💡 Clinical Pearl: In PRVC, a rising pressure trend over consecutive breaths is free information that compliance is deteriorating. Many people miss this.

[IMAGE: PRVC pressure waveform over 5-6 breaths showing the ventilator progressively increasing the pressure level as compliance worsens.]

# Section 4: Pressure-Volume Loops

The PV loop plots airway pressure (x-axis) against volume (y-axis) during a single breath, traced counterclockwise.

## 4.1 Normal PV Loop

A normal PV loop is an elongated oval. The slope represents dynamic compliance — steeper means more compliant. The area between limbs (hysteresis) represents energy lost to resistance and tissue friction.

[IMAGE: Normal PV loop labeled: inspiratory limb (right going up), expiratory limb (left coming down), compliance slope, hysteresis area shaded, starting point at PEEP.]

## 4.2 Lower Inflection Point (LIP)

The LIP is where compliance suddenly increases on the inspiratory limb — the slope steepens. Below this point, significant lung is collapsed. Historically used to guide PEEP setting, though modern strategies (ARDSNet tables, driving pressure) have largely replaced this approach.

[IMAGE: PV loop in ARDS showing a clear lower inflection point. Mark LIP with arrow, label optimal PEEP zone, show flat initial portion as atelectatic lung.]

## 4.3 Upper Inflection Point — Beaking

The UIP is where the inspiratory limb flattens at the top — more pressure yields progressively less volume. This creates a characteristic beak shape indicating overdistension and VILI risk.

🚨 OVERDISTENSION = LUNG INJURY RISK. If you see beaking: reduce tidal volume, reduce inspiratory pressure (ACPC), reassess driving pressure (target <15 cmH₂O).

[IMAGE: PV loop showing beaking — top of loop flattens and bends right. Label UIP, beak, overdistension zone. Overlay normal loop for comparison.]

## 4.4 Compliance Changes

| Change | PV Loop Appearance | Examples |
| --- | --- | --- |
| Decreased compliance | Loop shifts right, becomes flatter/wider | ARDS, pulmonary edema, PTX, abdominal compartment syndrome |
| Increased compliance | Loop shifts left, becomes steeper/narrower | Emphysema, post-thoracentesis, resolving edema |
| Increased resistance | Loop widens — larger hysteresis area | Bronchospasm, secretions, small ETT |
| Air trapping | Loop doesn’t close — expiratory limb incomplete | Auto-PEEP, inadequate expiratory time |

[IMAGE: Four PV loops overlaid: Normal (green), Decreased compliance (shifted right), Increased compliance (shifted left), Increased resistance (wider). Label each.]

## 4.5 Work of Breathing

The area to the left of the inspiratory limb represents work of breathing — shared between ventilator and patient. Anything increasing convexity or shifting the loop rightward increases total work of breathing.

[IMAGE: Two PV loops overlaid: Normal work of breathing (blue shading) vs severe asthma (red shading, rightward shift, more convex).]

## 4.6 Dyssynchrony on the PV Loop

- Tight trigger / auto-triggering: small loop-within-a-loop at bottom of PV curve

- Mid-inspiratory effort: leftward dip in inspiratory limb

- Flow starvation: inspiratory limb concave to the left

- Active exhalation against vent: rightward bulge on expiratory limb

[IMAGE: PV loop showing multiple dyssynchrony features on single breath — label each: trigger effort, mid-inspiratory dip, flow starvation concavity, active exhalation bulge.]

## 4.7 Air Leak on the PV Loop

When there is a circuit leak, the PV loop fails to close — the gap between start and end represents the lost volume.

[IMAGE: PV loop with circuit leak — loop doesn’t close. Label leak volume gap.]

# Section 5: Flow-Volume Loops

## 5.1 Normal FV Loop

Inspiration is the upper portion (positive flow), expiration is lower (negative flow). In ACVC with square flow, the inspiratory portion is flat. Expiration shows peak flow followed by decelerating curve.

[IMAGE: Normal FV loops: ACVC with square inspiratory flow and ACPC with decelerating inspiratory flow. Label inspiration and expiration.]

## 5.2 Obstructive Airway Disease

The expiratory limb shows a concave scooped appearance — peak expiratory flow drops rapidly then flow is prolonged and reduced. This is the FV equivalent of obstructive physiology.

[IMAGE: FV loop showing obstructive pattern — scooped expiratory limb vs normal overlay.]

## 5.3 Restrictive Lung Disease

The loop appears small and compressed with rapid decrease in inspiratory flow and high peak expiratory flow with rapid decay. Obstruction scoops the expiratory limb; restriction shrinks the entire loop.

[IMAGE: Three FV loops overlaid: Normal, Obstructive (scooped expiratory), Restrictive (small/compressed). Label each.]

## 5.4 Bronchodilator Response

After bronchodilators, the expiratory limb should show improved peak expiratory flow and a less concave curve. Comparing pre- and post-treatment FV loops at bedside is an excellent real-time assessment.

[IMAGE: Pre- and post-bronchodilator FV loops overlaid showing improvement in expiratory limb.]

## 5.5 Air Trapping on the FV Loop

If expiratory flow doesn’t reach zero before the next breath, the loop doesn’t close — visual confirmation of auto-PEEP.

[IMAGE: FV loop showing air trapping — expiratory flow cut off before reaching zero.]

# Section 6: Recognizing Dyssynchrony on Waveforms

Patient-ventilator dyssynchrony affects 25–50% of mechanically ventilated patients and is associated with prolonged ventilation and worse outcomes. The waveform is your first detection tool.

## 6.1 Flow Starvation

What it is: Patient’s inspiratory demand exceeds the ventilator’s set flow rate (ACVC).

Waveform: Pressure-time shows concave (scooped) rise during inspiration. Flow appears capped at set rate.

Fix: Increase peak flow to 60–80 L/min or higher, switch to decelerating flow, or convert to ACPC.

[IMAGE: Side-by-side pressure waveforms: (1) Normal convex rise, (2) Scooped/concave rise indicating flow starvation.]

## 6.2 Double Triggering

What it is: Single patient effort triggers two consecutive breaths because neural inspiratory time exceeds set Ti.

Waveform: Two complete inspiratory cycles back-to-back with no expiratory pause. Volume shows stacked 2× Vt.

Fix: Increase Ti, increase Vt slightly (if safe), switch to ACPC, or deepen sedation.

[IMAGE: Flow-time and volume-time showing double triggering — two breaths without expiratory pause, stacked volume labeled.]

## 6.3 Ineffective Triggering

What it is: Patient makes inspiratory effort that fails to trigger the ventilator. Effort is wasted.

Waveform: Small flow deflections during expiration that don’t trigger a breath. Small pressure dips during expiration.

Fix: Increase trigger sensitivity, address auto-PEEP, consider flow triggering.

[IMAGE: Flow-time waveform showing ineffective efforts — small upward deflections in expiratory flow that don’t trigger. Circle each missed effort.]

## 6.4 Auto-Triggering

What it is: Ventilator triggers a breath without patient effort — false triggering.

Waveform: No negative deflection before the breath. Caused by cardiac oscillations, circuit leaks, water.

Fix: Decrease trigger sensitivity, check for circuit leaks/water.

## 6.5 Premature Cycling

What it is: Ventilator ends inspiration before patient’s neural inspiration is complete.

Waveform: Pressure bump at inspiration-to-expiration transition.

Fix: Increase Ti, lower cycle-off threshold in PSV (25% → 10–15%).

## 6.6 Delayed Cycling

What it is: Ventilator continues delivering breath after neural inspiration ends.

Waveform: Sharp pressure spike at end-inspiration as patient pushes against continued flow.

Fix: Shorten Ti, increase cycle-off threshold in PSV (to 40–50%).

## 6.7 Reverse Triggering

What it is: Reflexive diaphragm contraction triggered by passive lung inflation. Seen in deep sedation.

Waveform: Small notch or dip in pressure plateau during mandatory breath.

Fix: Deepen sedation, re-paralyze briefly, consider APRV.

[IMAGE: Pressure waveform showing reverse triggering — notch in inspiratory plateau.]

# Section 7: Clinical Troubleshooting with Waveforms

## The 5-Step Systematic Approach

- Step 1: Check the baseline — Is PEEP where it should be?

- Step 2: Examine the inspiratory pressure waveform — Is PIP appropriate? Normal shape?

- Step 3: Look at the expiratory flow — Does it return to zero? Saw-tooth?

- Step 4: Compare inspired vs expired volumes — Any leak?

- Step 5: Assess the loops — Any beaking? Loop shape changing?

## Sudden High-Pressure Alarm (ACVC)

| Waveform Clue | Likely Cause | Immediate Action |
| --- | --- | --- |
| High PIP, normal Pplat | Airway resistance — secretions, bronchospasm, kinked ETT | Suction, bronchodilators, check tube |
| High PIP, high Pplat | Compliance — PTX, mainstem, edema, abdominal distension | CXR, assess abdomen, auscultate |
| High PIP, no plateau obtainable | Massive leak or circuit disconnect | Check all connections and cuff |

# Section 8: Auto-PEEP Detection & Quantification

Gold standard waveform sign: Expiratory flow on the flow-time scalar does not return to zero before the next breath.

[IMAGE: Flow-time scalar showing three consecutive breaths. First: normal. Second/third: expiratory flow cut off before zero. Zoomed inset.]

## Measuring Auto-PEEP

Expiratory hold maneuver: Hold expiratory hold button for 2–3 seconds. Auto-PEEP = Total PEEP − Set PEEP.

⚠️ The expiratory hold only works accurately in passive patients. In spontaneous breathing, the waveform sign (flow not reaching zero) is your primary detection tool.

## Consequences in Heart Failure

- Acts as unintended PEEP — increases intrathoracic pressure, reduces venous return

- In RV failure: can cause acute hemodynamic decompensation by increasing RV afterload

- In LVAD patients: reduces RV preload → decreased LVAD preload → low flows and suction events

- Causes triggering failure — patient must overcome auto-PEEP + trigger sensitivity

# Section 9: Waveform Changes with Interventions

## 9.1 Bronchodilator Response

Before: High PIP with normal Pplat, saw-tooth expiratory flow, scooped FV loop

After: PIP decreases, PIP-Pplat gap narrows, expiratory flow returns to zero faster, FV loop fuller

## 9.2 Suctioning Response

Before: Saw-tooth expiratory flow pattern, possibly elevated PIP

After: Flow smooths out, PIP decreases, Pplat unchanged (confirming resistive problem)

## 9.3 PEEP Titration

Recruitment: Compliance improves → PV loop shifts left, driving pressure may decrease

Overdistension: Beaking appears on PV loop, Pplat rises, Vt may drop (ACPC)

## 9.4 Diuresis in Pulmonary Edema

Progressive improvement over hours: Pplat decreases, PV loop shifts left and steepens, beaking resolves.

[IMAGE: Sequential PV loops showing improvement with diuresis — loops shift left and steepen. Label pre- and post-diuresis.]

## 9.5 Paralytic Administration

After paralysis: completely uniform waveforms. Ideal time to accurately measure Pplat, compliance, resistance.

💡 Clinical Pearl: If you need accurate lung mechanics, paralyze the patient. Spontaneous effort makes Pplat unreliable. A single dose of NMB for 2–3 minutes is sufficient.

# Section 10: Heart Failure & MCS-Specific Waveform Pearls

## 10.1 Cardiogenic Pulmonary Edema

Progressively rising Pplat over hours (worsening compliance) with normal PIP-Pplat gap. PV loop shifts right. Treatment response: diuresis should show progressive improvement on waveforms.

## 10.2 RV Failure & Pulmonary Hypertension

Every positive pressure breath increases RV afterload. Keep mean airway pressure as low as possible. Auto-PEEP is especially dangerous — monitor flow-time continuously.

## 10.3 LVAD Patients

Mean airway pressure directly affects RV preload and LVAD flows. If rising PIP or auto-PEEP AND falling LVAD flows — the ventilator is the culprit.

## 10.4 VA-ECMO Lung Rest

Ultra-protective settings: ACPC with low driving pressure, low RR, moderate PEEP. Watch for compliance improving (Vt rising at same pressure) — reduce pressure to maintain ultra-protective volumes.

## 10.5 Post-Cardiac Surgery

- Phrenic nerve injury: unilateral diaphragm paralysis → fails SBTs despite normal-appearing mechanics

- Tamponade: acute compliance drop with rising Pplat + hemodynamic instability

- Chest tube air leak: exhaled volume < inspired volume on volume-time scalar — quantify over time

# Section 11: Rapid Reference Tables

## 11.1 The Waveform Told Me — Clinical Decision Cheat Sheet

| You See This | It Means | You Do This |
| --- | --- | --- |
| PIP ↑, Pplat normal | Airway resistance ↑ | Suction, bronchodilators, check ETT |
| PIP ↑, Pplat ↑ | Compliance ↓ | CXR, rule out PTX/edema/effusion |
| Expiratory flow ≠ 0 | Auto-PEEP | Extend expiratory time, ↓ RR |
| Saw-tooth expiratory flow | Secretions | Suction |
| Scooped inspiratory pressure | Flow starvation | ↑ Flow or switch to ACPC |
| PV loop beaking | Overdistension | ↓ Vt, ↓ pressure, reassess PEEP |
| PV loop not closing | Leak or air trapping | Check cuff/circuit; assess auto-PEEP |
| FV expiratory limb scooped | Airway obstruction | Bronchodilators, assess on loop |
| Exhaled < inspired volume | Circuit or cuff leak | Check cuff pressure, connections |
| Pplat trending up over hours | Worsening compliance | Investigate — edema, ARDS, effusion |
| Two breaths back-to-back | Double triggering | ↑ Ti, consider mode change |
| Small flow blips, no trigger | Missed/ineffective efforts | ↑ Trigger sensitivity, treat auto-PEEP |

## 11.2 Dyssynchrony Quick Reference

| Dyssynchrony | Key Waveform Sign | Primary Fix |
| --- | --- | --- |
| Flow starvation | Scooped pressure during inspiration (ACVC) | ↑ Flow rate or switch to ACPC |
| Double triggering | Two breaths without expiratory pause | ↑ Ti, switch to ACPC, ↑ sedation |
| Ineffective triggering | Small flow/pressure deflections that don’t trigger | ↑ Trigger sensitivity, treat auto-PEEP |
| Auto-triggering | Breaths fire without patient effort | ↓ Trigger sensitivity, check circuit |
| Premature cycling | Pressure bump at insp→exp transition | ↑ Ti, ↓ cycle-off threshold (PSV) |
| Delayed cycling | Pressure spike at end-inspiration | ↓ Ti, ↑ cycle-off threshold (PSV) |
| Reverse triggering | Notch in pressure plateau during mandatory breath | Deepen sedation, re-paralyze |

# References & Further Reading

- Mellott KG, et al. The Basics of Ventilator Waveforms. Neonatal Netw. 2021;40(1):51-62.

- Oto B, et al. Patient-ventilator dyssynchrony in the ICU: A practical approach. Anaesth Intensive Care. 2021;49(2):86-97.

- Hess DR. Respiratory Mechanics in Mechanically Ventilated Patients. Respir Care. 2014;59(11):1773-1794.

- Harris RS. Pressure-Volume Curves of the Respiratory System. Respir Care. 2005;50(1):78-98.

- Yartsev A. Ventilator Waveform Analysis. Deranged Physiology (derangedphysiology.com).

- LITFL — Ventilator Waveform Analysis, PV Loops, Patient-Ventilator Dyssynchrony (litfl.com).

This document is a standalone reference for ventilator waveform interpretation. For mode-specific management protocols, see protocols-vent-management.md and protocols-vent-management-advanced.md.